## Shree Scope
Shree Calibration | Scope | ISO/IEC 17025 : 2017
Discipline / Group | Measurand or Reference Material/Type of instrument or material to be calibrated or measured / Quantity Measured /Instrument | Calibration or Measurement Method or procedure | Measurement range and additional parameters where applicable.
At Lab
MECHANICAL- DIMENSION (BASIC MEASURING INSTRUMENT, GAUGE ETC.) | Bore Gauge (For transmission) | Comparison by dial calibration tester | 0 to 1 mm
| Caliper (Vernier, Dial & Digital)(L.C. 0.01mm & coarser) | Comparison by using Caliper checker & slip gauge block | 0 to 600 mm
| Coating Thickness Meter (L.C. 0.001 mm) | Comparison by Thickness foils | 0 to 1 mm
| Comparator Stand For Flatness of Base | Comparison by Electronic comparator | 0 to 100 mm
| Comparator Stand For Flatness of Base | Comparison by Electronic comparator | 0 to 300 mm
| Depth Gauge (Vernier, Dial & Digital)(L.C. 0.01mm & coarser) | Comparison by Using Depth Micrometer Checker | 0 to 300mm
MECHANICAL- DIMENSION (BASIC MEASURING INSTRUMENT, GAUGE ETC.) | Depth Micrometer (L.C. 0.01mm & coarser) | Comparison by Depth Micrometer Checker | 0 to 300 mm
| Dial Calibration Tester (L.C. 0.0001mm) | Comparison by slip gauge block | 0 to 25 mm
| Dial Snap Gauge | Comparison by slip gauges | 1 mm to 200 mm
| Dial Thickness Gauge (L.C. 0.001 mm & coarser) | Comparison by Slip Gauge | 0 to 10 mm
| Electronic Probe With Indicator(L.C. 0.0001mm) | Comparison by Slip Gauge | 0 to 25 mm
| External Micrometer (L.C. 0.001mm & coarser) | Comparison by using Micrometer Checking set & Long Slip Gauge Block | 0 to 100 mm
| External Micrometer (L.C. 0.001mm & coarser) | Comparison by using Micrometer Checking set & Long Slip Gauge Block | 100 mm to 300 mm
| Feeler Gauge / Thickness foil | Comparison by electronic comparator | 0 to 1 mm
| Height Gauge (Vernier, Dial & Digital)(L.C. 0.01mm & coarser) | Comparison by using Caliper checker & Slip Gauge Block | 0 to 600 mm
| Inside Dial Caliper | Comparison by Slip gauges | 0 to 50 mm
MECHANICAL- DIMENSION (BASIC MEASURING INSTRUMENT, GAUGE ETC.) | Lever Dial (L.C. 0.001mm) | Comparison by dial calibration tester | 0 to 1 mm
| Measuring Pin | Comparison by Electronic Comparator & Slip Gauge | 0.5 mm to 25 mm
| Micrometer Setting Standard | Comparison by Electronic Comparator & Slip Gauge | 25 mm to 300mm
| Parallel Thread Plug Gauge | Comparison by FCDM,CSM & thread measuring wire | 2 mm to 300 mm
| Pistol Caliper (L.C. 0.10 mm) | Comparison by Slip Gauge | 0 to 50 mm
| Plain Plug Gauge / Width Gauge / Height Master / O.D. Gauge | Comparison by Electronic Comparator & Slip Gauge | 0 to 200 mm
| Plain Plug Gauge / Width Gauge / Height Master / O.D. Gauge | Comparison by Electronic Comparator & Slip Gauge | 200 mm to 400 mm
| Plain Snap Gauge | Comparison by slip gauges | 0.5 mm to 200 mm
| Plain Snap Gauge | Comparison by slip gauges | 200 mm to 400 mm
| Plunger Dial (L.C. 0.001mm & coarser) | Comparison by using dial calibration tester | 0 to 25 mm
| Taper Thread Plug Gauge | Comparison by FCDM,CSM & thread measuring wire | 2 mm to 100 mm
| Thread measuring wire | Comparison by Electronic Comparator & Slip Gauge | 0.17 mm to 6.35 mm
MECHANICAL- PRESSURE INDICATING DEVICES | Hydraulic Pressure - (Dial, Digital, Analogue gauges) | Comparison by Digital Pressure Gauge with hydraulic Test Pump as per DKD R6–1 | 0 to 70 bar
MECHANICAL- PRESSURE INDICATING DEVICES | Hydraulic Pressure - (Dial, Digital, Analogue gauges) | Comparison by Digital Pressure Gauge with hydraulic Test Pump as per DKD R6–1 | 0 to 700 bar
MECHANICAL- PRESSURE INDICATING DEVICES | Pneumatic Pressure - (Dial, Digital, Analogue gauges) | Comparison by Digital Pressure Gauge with pneumatic Test Pump as per DKD R6–1 | 0 to 7 bar
MECHANICAL- PRESSURE INDICATING DEVICES | Vacuum Pressure - (Dial, Digital, Analogue gauges) | Comparison by Digital Pressure Gauge with pneumatic Test Pump as per DKD R6–1 | -0.8 bar to 0 bar
MECHANICAL- TORQUE GENERATING DEVICES | Torque Wrench (Type I/Class B,C and Type II/Class A,B) Clockwise direction only | Comparison by torque transducer with indicator , IS16906:2018 | 2 Nm to 1000 Nm
At Site
MECHANICAL- DIMENSION (BASIC MEASURING INSTRUMENT, GAUGE ETC.) | Electronic Height Gauge(L.C. 0.0001 mm) | Comparison by long slip gauge blocks | 0 to 600 mm
| Surface Plate (Cast Iron & Granite) | Comparison by Spirit Level of sens. 0.01 mm/meter | 3000 mm to 3000 mm
MECHANICAL- PRESSURE INDICATING DEVICES | Hydraulic Pressure - (Dial, Digital, Analogue gauges) | Comparison by Digital Pressure Gauge with hydraulic Test Pump as per DKD R6–1 | 0 to 70 bar
| Hydraulic Pressure - (Dial, Digital, Analogue gauges) | Comparison by Digital Pressure Gauge with hydraulic Test Pump as per DKD R6–1 | 0 to 700 bar
| Pneumatic Pressure - (Dial, Digital, Analogue gauges) | Comparison by Digital Pressure Gauge with pneumatic Test Pump as per DKD R6–1 | 0 to 7 bar
| Vacuum Pressure - (Dial, Digital, Analogue gauges) | Comparison by Digital Pressure Gauge with pneumatic Test Pump as per DKD R6–1 | -0.8 bar to 0 bar
Thermal | Temp. indicator / controller | uni calibrator
| Glass Thermometer | wet bath, temp sensors
| Oven / furnace / autoclave / incubator / hot plate / water bath / Salt spray tester / Climatic chamber / Refrigerator / Deep freezer | temp sensors
Thermal Temperature | Liquid in glass/Dial/Digital Thermometer/Temperature Gauge | Using Standard PT-100 Sensor With 6 1/2 DMM, by Comparison Method | 30 °C to 250°C | 0.2°C
| RTD Sensor ,Thermocouple,Temperature Sensor ,With And Without Indicator,Data Logger And Recorder,Temperature Transmitter | Using Standard PT-100 Sensor With 6 1/2 DMM, oil bath by Comparison Method | 30°C to 250°C | 0.2°C
| Temperature indicator of freezer, Incuberator,Oven(At Single Point) | Using RTD And DMM | -40°C to 250°C | 3.9°C
| Temperature indicator Furnace(At Single Point) | Using R Type Thermocouple and 6 1/2 DMM | 300°C to 1100°C | 4.3°C
| Thermocouple with & Without indicator,Dataloogger and Recorder, Temperature Transmitter | Using R-Type Thermocouple,Dry Bath and 6 1/2DMM by Comparison Method | 300°C to 1100°C | 2.3°C
Thermal -Specific Heat And Humidity | Relative Humidity Sensor With Indicator,Digital Analog Thermohygrometer Humidity Chamber,Dew Point Meter | Using Standard RH Indicator With Sensor, Humidity Generated By Comparison Method | 20°C to 45 °C at 50% RH | 0.7°C
| Relative Humidity Sensor With Indicator,Digital Analog Thermohygrometer Humidity Chamber | Using Standard RH Indicator With Sensor, Humidity Generated By Comparison Method | 30% to 90%RH at 25°C | 2.1%
Electro Technical -Alternating Current (<1 GHz)(Measure) | AC Current At 50Hz | Using 6 1/2 DMM By Direct Method | 0.1 A to 1 A | 0.009% to 0.12%
| AC Current At 50Hz | Using 6 1/2 DMM By Direct Method | 1 A to 10 A | 0.12% to 0.10% | Oven / furnace / autoclave / incubator / hot plate / water bath / Salt spray tester / Climatic chamber / Refrigerator / Deep freezer
| AC Current At 50Hz | Using 6 1/2 DMM By Direct Method | 1 mA to 10 mA | 0.06% to 0.024%
| AC Current At 50Hz | Using 6 1/2 DMM By Direct Method | 10 mA to 100 mA | 0.024% to 0.009% | temp sensors
Electro Technical -Alternating Current (<1 GHz)(Measure) | AC Voltage At 50Hz | Using 6 1/2 DMM By Direct Method | 0.1 V to 1 V | 0.06% to 0.51%
| AC Voltage At 50Hz | Using 6 1/2 DMM By Direct Method | 1 V to 10 V | 0.51% to 0.04%
| AC Voltage At 50Hz | Using 6 1/2 DMM By Direct Method | 10 V to 100 V | 0.04% to 0.04%
| AC Voltage At 50Hz | Using 6 1/2 DMM By Direct Method | 100 V to 1000 V | 0.04% to 0.03%
| AC Voltage At 50Hz | Using 6 1/2 DMM By Direct Method | 2 mV to 100 mV | 0.1% t o 0.06%
Electro Technical -Alternating Current (<1 GHz)(Source) | AC Current At 50Hz | Using Multifunction Calibrator by Direct Method | 1 mA to 20 mA | 1.51% to 0.26%
| AC Current At 50Hz | Using Multifunction Calibrator With Current by Direct Method | 10 A to 1000A | 6.09% to 1.62%
| AC Current At 50Hz | Using Multifunction Calibrator by Direct Method | 2 A to 10 A | 0.39% to 0.32%
| AC Current At 50Hz | Using Multifunction Calibrator by Direct Method | 20 mA to 2000mA | 0.26% to 0.39%
Electro Technical -Alternating Current (<1 GHz)(Source) | AC Voltage At 50Hz | Using Multifunction Calibrator by Direct Method | 0.2 V to 2 V | 0.22% to 0.33%
| AC Voltage At 50Hz | Using Multifunction Calibrator by Direct Method | 2 mV to 200 mV | 1.12% to 0.22%
| AC Voltage At 50Hz | Using Multifunction Calibrator by Direct Method | 2 V to 900 V | 0.33% to 0.26%
Electro Technical -Direct Current (Measure) | DC Current | Using 6 1/2 DMM By Direct Method | 0.1 A to 1 A | 0.12% to 0.03%
| DC Current | Using 6 1/2 DMM By Direct Method | 1 A to 10 A | 0.03% to 0.1%
| DC Current | Using 6 1/2 DMM By Direct Method | 1 mA to 10 mA | 0.18% to 0.024%
| DC Current | Using 6 1/2 DMM By Direct Method | 10 mA to 100 mA | 0.024% to 0.009%
Electro Technical -Direct Current (Measure) | DC Voltage | Using 6 1/2 DMM By Direct Method | 0.1 V to 1 V | 0.01% to 0.03%
| DC Voltage | Using 6 1/2 DMM By Direct Method | 1 V to 10 V | 0.03% to 0.02%
| DC Voltage | Using 6 1/2 DMM By Direct Method | 10 V to 100 V | 0.02% to 0.001%
| DC Voltage | Using 6 1/2 DMM By Direct Method | 100 V to 1000 V | 0.01% to 0.035%
| DC Voltage | Using 6 1/2 DMM By Direct Method | 1 mV to 100 mV | 0.44% to 0.01%
Electro Technical -Direct Current (Source) | DC Current | Using Multifunction Calibrator by Direct Method | 1 mA to 20 mA | 1.48% to 0.16%
| DC Current | Using Multifunction Calibrator by Direct Method | 10 A to 1000A | 5.97% to 1.27%
| DC Current | Using Multifunction Calibrator by Direct Method | 2 A to 10 A | 0.22% to 0.18%
| DC Current | Using Multifunction Calibrator by Direct Method | 20 mA to 2000mA | 0.22%
Electro Technical -Direct Current (Source) | DC Voltage | Using Multifunction Calibrator by Direct Method | 0.2 V to 2 V | 0.22% to 0.16%
| DC Voltage | Using Multifunction Calibrator by Direct Method | 1 mV to 200 mV | 1.44% to 0.22%
| DC Voltage | Using Multifunction Calibrator by Direct Method | 2 V to 1000 V | 0.22% to 0.17%
Electro Technical -Temperature Simulation (Source) | Thermocouple J- Type (Temperature indicator,controller,Recorder | By Direct Method Using Universal Calibrator | 50°C to 600 °C | 0.84°C
| Thermocouple K- Type (Temperature indicator,controller,Recorder | By Direct Method Using Universal Calibrator | 50°C to 1200°C | 0.84°C
| Thermocouple R- Type (Temperature indicator,controller,Recorder | By Direct Method Using Universal Calibrator | 200°C to 1750°C | 1.42°C
| Thermocouple S- Type (Temperature indicator,controller,Recorder | By Direct Method Using Universal Calibrator | 50°C to 1750°C | 1.47°C
| Thermocouple T- Type (Temperature indicator,controller,Recorder | By Direct Method Using Universal Calibrator | -160°C to 400°C | 0.83°C
Electro Technical -Time and Frequency (Source) | Frequency | Using 6 1/2 DMM By Direct Method | 45000 Hz to 1000Hz | 1.28% to 0.06%
| Timer / Stop Watch | Using Programmable Timer By Comparison Metod | 4 s to 5 Hr | 0.02 s to 6.53 s
Fluid Flow-Flow Measuring Devices | Volumetric Flow Rate | Using Ultrasonic Flow Meter using HM Sensor Suitable 2 to 28 inch dia pipe by comparison method | 30 m3/hr to 300 m3/hr | 2.15% to 0.84%
Mechanical- Acceleration And Speed | RPM Meter/Speed Centrifuge(contact) | Using Digital Techometer By comparison Method | 10 to 5000rpm | 4rpm

## Master Equipment
SR.NO. | EQUIPMENT NAME | MAKE | RANGE/SIZE | LC/GRADE | CALIBRATION FREQUENCY |  |  |  | Status as On 08.01.2023
[6 empty] | Cost | Remark | Remark
1 | TEMP. AND HUMIDITY DATA LOGGER | -- | 0-100° C & 10-99% RH | 0.10° C 0.1%RH | 1 YEAR | 15000 | Testo |  | Orderd To Calitech
2 | MIC CHECK SET (10 PIECES) | kcp | 2.5-25mm | '0' GRADE | 2 YEARS | 19500 |  | Phase 2
3 | CALIPER CHECKER | kcp | 0 - 600 mm | --- | 2 YEARS | 54900 | Pawar |  | To be Order
4 | DIAL CALIBRATION TESTER | MIKROTECH | 0-10mm | 0.0001mm | 1 YEAR | 302700 | Pawar |  | Hold
5 | EXT MICROMETER (DIGITAL) | MITUTOYO | 0 - 25 mm | 0.001mm | 1 YEAR | 6950 | Order |  | Ordered to Hemant Tools
6 | PLUNGER TYPE DIAL GAUGE | MITUTOYO | 0 - 1 mm | 0.001 mm | 1 YEAR | 3550 | Order |  | Ordered to Hemant Tools
7 | SURFACE PLATE - GRANITE | JAFUGI | 610 x 610 mm | '0' GRADE | 2 YEARS |  | Pawar |  | To be Order Prcecision
8 | DIGITAL THERMOMETER | -- | -50-300° C | 0.10° C | 1 YEAR | 1000 | Testo |  | Orderd To Calitech
9 | PRECISION SPIRIT LEVEL | Millard | 200mm | 0.01mm/meter | 2 YEARS | 20000 | Order |  | Orderd To Calitech
10 | LONG SLIP GAUGES BLOCKS | kcp | 50-400mm | '0' GRADE | 2 YEARS | 83900 |  | Phase 2 | Hold
11 | ILLUMINATION & NOISE LEVEL TEST | -- | -- | -- | 2 YEARS | 2000 |  | Phase 2 | Hold
12 | DEPTH MICRO CHECKER | kcp | 0-300mm | -- | 2 YEARS | 57900 | Order |  | Ordered kcp
13 | SLIP GAUGE SET (87 PIECES) | kcp | 0.5-100mm | '0' GRADE | 2 YEARS | 109000 | Order |  | Ordered kcp
14 | SLIP GAUGE BLOCKS (2 PCS) | kcp | 10 mm | '0' GRADE | 2 YEARS | 3720 | Order | with Access | Hold
15 | FOILS | -- | 10-710μm | -- | 1 YEAR | 1000 | Order |  | Find Th Qautation
16 | ELECTRONIC PROBE WITH DRO | Varsa | 0-4 mm | 0.0001 mm | 2 YEARS | 65000 |  | Caltech | Orderd To Calitech
17 | DIGITAL PRESSURE GAUGE | Adarsh | -1-35 bar | 0.0001 bar | 1 YEAR | 30000 | Order for -1 to 35 bar | Pnematic & vaccum | Orderd To Adarsh
18 | DIGITAL PRESSURE GAUGE | Adarsh | 0-70 bar | 0.001 bar | 1 YEAR | 30000 | order | For hy | Orderd To Adarsh
19 | DIGITAL PRESSURE GAUGE | Adarsh | 0-700 bar | 0.01 bar | 1 YEAR | 31000 | order | For hy | Orderd To Adarsh
20 | Pnematic pump/ vaccum |  |  |  |  | 28000 | order |  | Orderd To Adarsh
21 | Hy Pump |  |  |  |  | 43500 | order |  | Orderd To Adarsh
22 | Secondary Pr Gauge |  | (-)1 to 50 bar |  |  | 8500 |  | phase 2 | Hold
23 | Secondary Pr Gauge |  | 0 to 700 bar |  |  | 10000 |  | phase 2 | Hold
24 | COMPARATOR STAND BASE | kcp | 95 X 90 mm, 0-1000 mm | -- | 2 YEARS | 22500 | Order | Ht should be 300mm | Ordered kcp
25 | TORQUE WRENCH TESTER | Suraksha sensors | 0-10 Nm 0-100 Nm 0-1000 Nm | 0.001 Nm 0.01 Nm 0.1 Nm | 1 YEAR | 235000 | Pawar |  | To be Disscused & Order
26 | FCDM | Aura/Kcp/Micronix | 0-100mm | 0.001mm | 1 YEAR |  |  | Phase 2 | Hold
27 | Cylindrical setting master | Kcp/micronix/jayesh | 5mm-100mm | - | 1 YEAR |  |  | Phase 2 | Hold
28 | Thread measuring wire | Kcp/micronix/jayesh | 0.17-6.23mm | - | 1 YEAR |  |  | Phase 2 | Hold
29 | universal Calibrator | Massibus |  |  | 1 YEAR | 75000 | Order | Caltech | Disscused
30 | wet bath With Temp sensors | Teschno | Ambent to 600 |  |  | 55000 | Order | Caltech | Disscused
31 | temp sensors | Teschno | PT 100 & thermocouple |  | 1 YEAR | 25000 | order | Caltech | Disscused

## Table 7
Shree Calibration Measurement And Testing Services
QUALITY MANUAL |  | ISSUE NO. / DATE | 01 / 01.11.2018
|  | REF. NO. | ANNEXURE– III
SCOPE OF ACCREDITATION |  | REV. NO. / DATE | 01 / 10.10.2022
|  | PAGE NO. | 1 of 5
Discipline / Group | Measurand or Reference Material/Type of instrument or material to be calibrated or measured / Quantity Measured /Instrument | Calibration or Measurement Method or procedure | Measurement range and additional parameters where applicable (Range and Frequency) | * Calibration and Measurement Capability (CMC)(±) | Scope | Master details | Master Supplier | Apprx Cost
At Lab
MECHANICAL- DIMENSION (BASIC MEASURING INSTRUMENT, GAUGE ETC.) | Bore Gauge (For transmission) | Comparison by dial calibration tester | 0 to 1 mm | 3.9µm | Y
| Caliper (Vernier, Dial & Digital)(L.C. 0.01mm & coarser) | Comparison by using Caliper checker & slip gauge block & Length Bar | 0 to 600 mm | 12.0µm | Y
| Coating Thickness Meter (L.C. 0.001 mm) | Comparison by Thickness foils | 0 to 1 mm | 8µm | Y
| Comparator Stand For Flatness of Base | Comparison by Electronic comparator | 0 to 100 mm | 1.6µm | Y
| Comparator Stand For Flatness of Base | Comparison by Electronic comparator | 0 to 300 mm | 3µm | Y
| Depth Gauge (Vernier, Dial & Digital)(L.C. 0.01mm & coarser) | Comparison by Using Depth Micrometer Checker | 0 to 300mm | 13.0µm | Y
PREPARED & ISSUED BY – TM |  | APPROVED BY - QM
CAL-TECH CALIBRATION SERVICES PVT. LTD.
QUALITY MANUAL |  | ISSUE NO. / DATE | 01 / 01.11.2018
|  | REF. NO. | ANNEXURE– III
SCOPE OF ACCREDITATION |  | REV. NO. / DATE | 01 / 10.10.2022
|  | PAGE NO. | 2 of 5
MECHANICAL- DIMENSION (BASIC MEASURING INSTRUMENT, GAUGE ETC.) | Depth Micrometer (L.C. 0.01mm & coarser) | Comparison by Depth Micrometer Checker | 0 to 300 mm | 12µm | Y
| Dial Calibration Tester (L.C. 0.0001mm) | Comparison by slip gauge block | 0 to 25 mm | 1.3µm | Y
| Dial Snap Gauge | Comparison by slip gauges | 1 mm to 200 mm | 3.5µm | Y
| Dial Thickness Gauge (L.C. 0.001 mm & coarser) | Comparison by Slip Gauge | 0 to 10 mm | 1µm | Y
| Electronic Probe With Indicator(L.C. 0.0001mm) | Comparison by Slip Gauge | 0 to 25 mm | 0.7µm | Y
| External Micrometer (L.C. 0.001mm & coarser) | Comparison by using Micrometer Checking set & Long Slip Gauge Block | 0 to 100 mm | 2µm | Y
| External Micrometer (L.C. 0.001mm & coarser) | Comparison by using Micrometer Checking set & Long Slip Gauge Block | 100 mm to 300 mm | 7µm | Y
| Feeler Gauge / Thickness foil | Comparison by electronic comparator | 0 to 1 mm | 1.3µm | Y
| Height Gauge (Vernier, Dial & Digital)(L.C. 0.01mm & coarser) | Comparison by using Caliper checker & Slip Gauge Block | 0 to 600 mm | 12.0µm | Y
| Inside Dial Caliper | Comparison by Slip gauges | 0 to 50 mm | 6.7µm | Y
| Internal Micrometer- 2 Point (for Micrometer head) (L.C. 0.01 mm) | Comparison by using ULM300, Slip Gauges blocks and comparator stand | 0 to 100 mm | 10µm | N
| Internal Micrometer- 2 Point (for Micrometer range with extension rods) (L.C. 0.01 mm) | Comparison by using ULM300, Slip Gauges blocks and comparator stand | 0 to 1000 mm | 12µm | N
PREPARED & ISSUED BY – TM |  | APPROVED BY - QM
CAL-TECH CALIBRATION SERVICES PVT. LTD.
QUALITY MANUAL |  | ISSUE NO. / DATE | 01 / 01.11.2018
|  | REF. NO. | ANNEXURE– III
SCOPE OF ACCREDITATION |  | REV. NO. / DATE | 01 / 10.10.2022
|  | PAGE NO. | 3 of 5
MECHANICAL- DIMENSION (BASIC MEASURING INSTRUMENT, GAUGE ETC.) | Lever Dial (L.C. 0.001mm) | Comparison by dial calibration tester | 0 to 1 mm | 2µm | Y
| Measuring Pin | Comparison by Electronic Comparator & Slip Gauge | 0.5 mm to 25 mm | 1.4µm | Y
| Measuring Scale | Comparison by Scale Calibrator | 0 to 2000 mm | 118*Sqrt(L), L in meterµm | N
| Measuring Tape / Pie Tape | Comparison by Scale Calibrator | 0 to 50 meter | 261*Sqrt(L), L in meterµm | N
| Micrometer Setting Standard | Comparison by Electronic Comparator & Slip Gauge | 25 mm to 300mm | 5.4µm | Y
| Parallel Thread Plug Gauge | Comparison by length measuring machine, thread measuring wire | 2 mm to 300 mm | 3µm | N
| Pistol Caliper (L.C. 0.10 mm) | Comparison by Slip Gauge | 0 to 50 mm | 66.7µm | Y
| Plain Plug Gauge / Width Gauge / Height Master / O.D. Gauge | Comparison by Electronic Comparator & Slip Gauge | 0 to 200 mm | 1.9µm | Y
| Plain Plug Gauge / Width Gauge / Height Master / O.D. Gauge | Comparison by Electronic Comparator & Slip Gauge | 200 mm to 400 mm | 3.5µm | Y
| Plain Snap Gauge | Comparison by slip gauges | 0.5 mm to 200 mm | 2.5µm | Y
| Plain Snap Gauge | Comparison by slip gauges | 200 mm to 400 mm | 3.3µm | Y
| Plunger Dial (L.C. 0.001mm & coarser) | Comparison by using dial calibration tester | 0 to 25 mm | 2µm | Y
PREPARED & ISSUED BY – TM |  | APPROVED BY - QM
CAL-TECH CALIBRATION SERVICES PVT. LTD.
QUALITY MANUAL |  | ISSUE NO. / DATE | 01 / 01.11.2018
|  | REF. NO. | ANNEXURE– III
SCOPE OF ACCREDITATION |  | REV. NO. / DATE | 01 / 10.10.2022
|  | PAGE NO. | 4 of 5
| Right Angle / Engineers Square - Squareness | Comparison by Cylindrical Square Master | 0 to 450 mm | 10µm | N
| Taper Thread Plug Gauge | Comparison by FCDM, thread measuring wire | 2 mm to 100 mm | 4µm | N
| Thread measuring wire | Comparison by Electronic Comparator & Slip Gauge | 0.17 mm to 6.35 mm | 1µm | N
| V Block - Parallelism | Comparison by Mandrel & Lever Dial | 0 to 200 mm | 15.3µm | N
| V Block - Squareness | Comparison by Cylindrical Square Master | 0 to 200 mm | 8.6µm | N
| V Block - Symmetricity | Comparison by Mandrel & Lever Dial | 0 to 200 mm | 15.3µm | N
MECHANICAL- PRESSURE INDICATING DEVICES | Hydraulic Pressure - (Dial, Digital, Analogue gauges) | Comparison by Digital Pressure Gauge with hydraulic Test Pump as per DKD R6–1 | 0 to 70 bar | 0.17bar | Y
MECHANICAL- PRESSURE INDICATING DEVICES | Hydraulic Pressure - (Dial, Digital, Analogue gauges) | Comparison by Digital Pressure Gauge with hydraulic Test Pump as per DKD R6–1 | 0 to 700 bar | 1.16bar | Y
PREPARED & ISSUED BY – TM |  | APPROVED BY - QM
CAL-TECH CALIBRATION SERVICES PVT. LTD.
QUALITY MANUAL |  | ISSUE NO. / DATE | 01 / 01.11.2018
|  | REF. NO. | ANNEXURE– III
SCOPE OF ACCREDITATION |  | REV. NO. / DATE | 01 / 10.10.2022
|  | PAGE NO. | 5 of 5
MECHANICAL- PRESSURE INDICATING DEVICES | Pneumatic Pressure - (Dial, Digital, Analogue gauges) | Comparison by Digital Pressure Gauge with pneumatic Test Pump as per DKD R6–1 | 0 to 7 bar | 0.06bar | Y
MECHANICAL- PRESSURE INDICATING DEVICES | Vacuum Pressure - (Dial, Digital, Analogue gauges) | Comparison by Digital Pressure Gauge with pneumatic Test Pump as per DKD R6–1 | -0.8 bar to 0 bar | 0.04bar | Y
MECHANICAL- TORQUE GENERATING DEVICES | Torque Wrench (Type I/Class B,C and Type II/Class A,B) Clockwise direction only | Comparison by torque transducer with indicator , IS16906:2018 | 2 Nm to 1000 Nm | 2% | Y
At Site
MECHANICAL- DIMENSION (BASIC MEASURING INSTRUMENT, GAUGE ETC.) | Electronic Height Gauge(L.C. 0.0001 mm) | Comparison by long slip gauge blocks | 0 to 600 mm | 6µm | Y
| Surface Plate (Cast Iron & Granite) | Comparison by Spirit Level of sens. 0.01 mm/meter | 3000 mm to 3000 mm | 4*sqrt(L+W/150), L in mmµm | Y
MECHANICAL- DIMENSION (PRECISION INSTRUMENTS) | Proﬁle Projector / Optical Microscope (Angular L.C. 1 second) | Comparison by Angle Gauge Blocks | 0 to 360 ° | 2min. | N
| Proﬁle Projector / Optical Microscope (Linear X,Y Axis L.C. 0.001mm) | Comparison by Precision Glass Scale | 0 to 200 mm | 6.2µm | N
| Proﬁle Projector / Optical Microscope (Magniﬁcation) | Comparison by Digital Caliper & Glass scale | 10X to 100X | 0.2% | N
MECHANICAL- PRESSURE INDICATING DEVICES | Hydraulic Pressure - (Dial, Digital, Analogue gauges) | Comparison by Digital Pressure Gauge with hydraulic Test Pump as per DKD R6–1 | 0 to 70 bar | 0.17bar | Y
| Hydraulic Pressure - (Dial, Digital, Analogue gauges) | Comparison by Digital Pressure Gauge with hydraulic Test Pump as per DKD R6–1 | 0 to 700 bar | 1.16bar | Y
| Pneumatic Pressure - (Dial, Digital, Analogue gauges) | Comparison by Digital Pressure Gauge with pneumatic Test Pump as per DKD R6–1 | 0 to 7 bar | 0.06bar | Y
| Vacuum Pressure - (Dial, Digital, Analogue gauges) | Comparison by Digital Pressure Gauge with pneumatic Test Pump as per DKD R6–1 | -0.8 bar to 0 bar | 0.04bar | Y
PREPARED & ISSUED BY – TM |  | APPROVED BY - QM

## Shree Scope With Std
Shree Calibration | Scope
Discipline / Group | Measurand or Reference Material/Type of instrument or material to be calibrated or measured / Quantity Measured /Instrument | Calibration or Measurement Method or procedure | Measurement range and additional parameters where applicable.
At Lab
MECHANICAL- DIMENSION (BASIC MEASURING INSTRUMENT, GAUGE ETC.) | Bore Gauge (For transmission) | Comparison by dial calibration tester | 0 to 1 mm | JIS B 7515
| Caliper (Vernier, Dial & Digital)(L.C. 0.01mm & coarser) | Comparison by using Caliper checker & slip gauge block | 0 to 600 mm | IS 3651(Part I,II,III)
| Coating Thickness Meter (L.C. 0.001 mm) | Comparison by Thickness foils | 0 to 1 mm
| Comparator Stand For Flatness of Base | Comparison by Electronic comparator | 0 to 100 mm
| Comparator Stand For Flatness of Base | Comparison by Electronic comparator | 0 to 300 mm
| Depth Gauge (Vernier, Dial & Digital)(L.C. 0.01mm & coarser) | Comparison by Using Depth Micrometer Checker | 0 to 300mm | IS 4213
MECHANICAL- DIMENSION (BASIC MEASURING INSTRUMENT, GAUGE ETC.) | Depth Micrometer (L.C. 0.01mm & coarser) | Comparison by Depth Micrometer Checker | 0 to 300 mm
| Dial Calibration Tester (L.C. 0.0001mm) | Comparison by slip gauge block | 0 to 25 mm
| Dial Snap Gauge | Comparison by slip gauges | 1 mm to 200 mm | IS 3477
| Dial Thickness Gauge (L.C. 0.001 mm & coarser) | Comparison by Slip Gauge | 0 to 10 mm
| Electronic Probe With Indicator(L.C. 0.0001mm) | Comparison by Slip Gauge | 0 to 25 mm
| External Micrometer (L.C. 0.001mm & coarser) | Comparison by using Micrometer Checking set & Long Slip Gauge Block | 0 to 100 mm | IS 2967
| External Micrometer (L.C. 0.001mm & coarser) | Comparison by using Micrometer Checking set & Long Slip Gauge Block | 100 mm to 300 mm
| Feeler Gauge / Thickness foil | Comparison by electronic comparator | 0 to 1 mm
| Height Gauge (Vernier, Dial & Digital)(L.C. 0.01mm & coarser) | Comparison by using Caliper checker & Slip Gauge Block | 0 to 600 mm | IS 2921
| Inside Dial Caliper | Comparison by Slip gauges | 0 to 50 mm
MECHANICAL- DIMENSION (BASIC MEASURING INSTRUMENT, GAUGE ETC.) | Lever Dial (L.C. 0.001mm) | Comparison by dial calibration tester | 0 to 1 mm | IS 11498
| Measuring Pin | Comparison by Electronic Comparator & Slip Gauge | 0.5 mm to 25 mm
| Micrometer Setting Standard | Comparison by Electronic Comparator & Slip Gauge | 25 mm to 300mm
| Parallel Thread Plug Gauge | Comparison by FCDM,CSM & thread measuring wire | 2 mm to 300 mm
| Pistol Caliper (L.C. 0.10 mm) | Comparison by Slip Gauge | 0 to 50 mm
| Plain Plug Gauge / Width Gauge / Height Master / O.D. Gauge | Comparison by Electronic Comparator & Slip Gauge | 0 to 200 mm
| Plain Plug Gauge / Width Gauge / Height Master / O.D. Gauge | Comparison by Electronic Comparator & Slip Gauge | 200 mm to 400 mm
| Plain Snap Gauge | Comparison by slip gauges | 0.5 mm to 200 mm
| Plain Snap Gauge | Comparison by slip gauges | 200 mm to 400 mm
| Plunger Dial (L.C. 0.001mm & coarser) | Comparison by using dial calibration tester | 0 to 25 mm | IS 2092
| Taper Thread Plug Gauge | Comparison by FCDM,CSM & thread measuring wire | 2 mm to 100 mm
| Thread measuring wire | Comparison by Electronic Comparator & Slip Gauge | 0.17 mm to 6.35 mm
MECHANICAL- PRESSURE INDICATING DEVICES | Hydraulic Pressure - (Dial, Digital, Analogue gauges) | Comparison by Digital Pressure Gauge with hydraulic Test Pump as per DKD R6–1 | 0 to 70 bar
MECHANICAL- PRESSURE INDICATING DEVICES | Hydraulic Pressure - (Dial, Digital, Analogue gauges) | Comparison by Digital Pressure Gauge with hydraulic Test Pump as per DKD R6–1 | 0 to 700 bar
MECHANICAL- PRESSURE INDICATING DEVICES | Pneumatic Pressure - (Dial, Digital, Analogue gauges) | Comparison by Digital Pressure Gauge with pneumatic Test Pump as per DKD R6–1 | 0 to 7 bar
MECHANICAL- PRESSURE INDICATING DEVICES | Vacuum Pressure - (Dial, Digital, Analogue gauges) | Comparison by Digital Pressure Gauge with pneumatic Test Pump as per DKD R6–1 | -0.8 bar to 0 bar
MECHANICAL- TORQUE GENERATING DEVICES | Torque Wrench (Type I/Class B,C and Type II/Class A,B) Clockwise direction only | Comparison by torque transducer with indicator , IS16906:2018 | 2 Nm to 1000 Nm
At Site
MECHANICAL- DIMENSION (BASIC MEASURING INSTRUMENT, GAUGE ETC.) | Electronic Height Gauge(L.C. 0.0001 mm) | Comparison by long slip gauge blocks | 0 to 600 mm
| Surface Plate (Cast Iron & Granite) | Comparison by Spirit Level of sens. 0.01 mm/meter | 3000 mm to 3000 mm | IS 2220,IS12937
MECHANICAL- PRESSURE INDICATING DEVICES | Hydraulic Pressure - (Dial, Digital, Analogue gauges) | Comparison by Digital Pressure Gauge with hydraulic Test Pump as per DKD R6–1 | 0 to 70 bar
| Hydraulic Pressure - (Dial, Digital, Analogue gauges) | Comparison by Digital Pressure Gauge with hydraulic Test Pump as per DKD R6–1 | 0 to 700 bar
| Pneumatic Pressure - (Dial, Digital, Analogue gauges) | Comparison by Digital Pressure Gauge with pneumatic Test Pump as per DKD R6–1 | 0 to 7 bar
| Vacuum Pressure - (Dial, Digital, Analogue gauges) | Comparison by Digital Pressure Gauge with pneumatic Test Pump as per DKD R6–1 | -0.8 bar to 0 bar
Thermal | Temp. indicator / controller | uni calibrator
| Glass Thermometer | wet bath, temp sensors
| Oven / furnace / autoclave / incubator / hot plate / water bath / Salt spray tester / Climatic chamber / Refrigerator / Deep freezer | temp sensors |  | IEC60068(Part 3-6),Part11,DKD-R5-7
Thermal Temperature | Liquid in glass/Dial/Digital Thermometer/Temperature Gauge | Using Standard PT-100 Sensor With 6 1/2 DMM, by Comparison Method | 30 °C to 250°C | IS 6274, IS 2480,OIL R 133
| RTD Sensor ,Thermocouple,Temperature Sensor ,With And Without Indicator,Data Logger And Recorder,Temperature Transmitter | Using Standard PT-100 Sensor With 6 1/2 DMM, oil bath by Comparison Method | 30°C to 250°C | DKD- R-5-1
| Temperature indicator of freezer, Incuberator,Oven(At Single Point) | Using RTD And DMM | -40°C to 250°C
| Temperature indicator Furnace(At Single Point) | Using R Type Thermocouple and 6 1/2 DMM | 300°C to 1100°C | ASTM E220-13/EURAMET cg-8
| Thermocouple with & Without indicator,Dataloogger and Recorder, Temperature Transmitter | Using R-Type Thermocouple,Dry Bath and 6 1/2DMM by Comparison Method | 300°C to 1100°C
Thermal -Specific Heat And Humidity | Relative Humidity Sensor With Indicator,Digital Analog Thermohygrometer Humidity Chamber | Using Standard RH Indicator With Sensor, Humidity Generated By Comparison Method | 20°C to 45 °C at 50% RH
| Relative Humidity Sensor With Indicator,Digital Analog Thermohygrometer Humidity Chamber | Using Standard RH Indicator With Sensor, Humidity Generated By Comparison Method | 30% to 90%RH at 25°C | DKD- R5-7
Electro Technical -Alternating Current (<1 GHz)(Measure) | AC Current At 50Hz | Using 6 1/2 DMM By Direct Method | 0.1 A to 1 A
| AC Current At 50Hz | Using 6 1/2 DMM By Direct Method | 1 A to 10 A |  | Oven / furnace / autoclave / incubator / hot plate / water bath / Salt spray tester / Climatic chamber / Refrigerator / Deep freezer
| AC Current At 50Hz | Using 6 1/2 DMM By Direct Method | 1 mA to 10 mA
| AC Current At 50Hz | Using 6 1/2 DMM By Direct Method | 10 mA to 100 mA |  | temp sensors
Electro Technical -Alternating Current (<1 GHz)(Measure) | AC Voltage At 50Hz | Using 6 1/2 DMM By Direct Method | 0.1 V to 1 V
| AC Voltage At 50Hz | Using 6 1/2 DMM By Direct Method | 1 V to 10 V
| AC Voltage At 50Hz | Using 6 1/2 DMM By Direct Method | 10 V to 100 V
| AC Voltage At 50Hz | Using 6 1/2 DMM By Direct Method | 100 V to 1000 V
| AC Voltage At 50Hz | Using 6 1/2 DMM By Direct Method | 2 mV to 100 mV
Electro Technical -Alternating Current (<1 GHz)(Source) | AC Current At 50Hz | Using Multifunction Calibrator by Direct Method | 1 mA to 20 mA
| AC Current At 50Hz | Using Multifunction Calibrator With Current by Direct Method | 10 A to 1000A
| AC Current At 50Hz | Using Multifunction Calibrator by Direct Method | 2 A to 10 A
| AC Current At 50Hz | Using Multifunction Calibrator by Direct Method | 20 mA to 2000mA
Electro Technical -Alternating Current (<1 GHz)(Source) | AC Voltage At 50Hz | Using Multifunction Calibrator by Direct Method | 0.2 V to 2 V
| AC Voltage At 50Hz | Using Multifunction Calibrator by Direct Method | 2 mV to 200 mV
| AC Voltage At 50Hz | Using Multifunction Calibrator by Direct Method | 2 V to 900 V
Electro Technical -Direct Current (Measure) | DC Current | Using 6 1/2 DMM By Direct Method | 0.1 A to 1 A
| DC Current | Using 6 1/2 DMM By Direct Method | 1 A to 10 A
| DC Current | Using 6 1/2 DMM By Direct Method | 1 mA to 10 mA
| DC Current | Using 6 1/2 DMM By Direct Method | 10 mA to 100 mA
Electro Technical -Direct Current (Measure) | DC Voltage | Using 6 1/2 DMM By Direct Method | 0.1 V to 1 V
| DC Voltage | Using 6 1/2 DMM By Direct Method | 1 V to 10 V
| DC Voltage | Using 6 1/2 DMM By Direct Method | 10 V to 100 V
| DC Voltage | Using 6 1/2 DMM By Direct Method | 100 V to 1000 V
| DC Voltage | Using 6 1/2 DMM By Direct Method | 1 mV to 100 mV
Electro Technical -Direct Current (Source) | DC Current | Using Multifunction Calibrator by Direct Method | 1 mA to 20 mA
| DC Current | Using Multifunction Calibrator by Direct Method | 10 A to 1000A
| DC Current | Using Multifunction Calibrator by Direct Method | 2 A to 10 A
| DC Current | Using Multifunction Calibrator by Direct Method | 20 mA to 2000mA
Electro Technical -Direct Current (Source) | DC Voltage | Using Multifunction Calibrator by Direct Method | 0.2 V to 2 V
| DC Voltage | Using Multifunction Calibrator by Direct Method | 1 mV to 200 mV
| DC Voltage | Using Multifunction Calibrator by Direct Method | 2 V to 1000 V
Electro Technical -Temperature Simulation (Source) | Thermocouple J- Type (Temperature indicator,controller,Recorder | By Direct Method Using Universal Calibrator | 50°C to 600 °C
| Thermocouple K- Type (Temperature indicator,controller,Recorder | By Direct Method Using Universal Calibrator | 50°C to 1200°C
| Thermocouple R- Type (Temperature indicator,controller,Recorder | By Direct Method Using Universal Calibrator | 200°C to 1750°C
| Thermocouple S- Type (Temperature indicator,controller,Recorder | By Direct Method Using Universal Calibrator | 50°C to 1750°C
| Thermocouple T- Type (Temperature indicator,controller,Recorder | By Direct Method Using Universal Calibrator | -160°C to 400°C
Electro Technical -Time and Frequency (Source) | Frequency | Using 6 1/2 DMM By Direct Method | 45000 Hz to 1000Hz
| Timer / Stop Watch | Using Programmable Timer By Comparison Metod | 4 s to 5 Hr

## scope with std no
Shree Calibration | Scope
Discipline / Group | Measurand or Reference Material/Type of instrument or material to be calibrated or measured / Quantity Measured /Instrument | Calibration or Measurement Method or procedure | Measurement range and additional parameters where applicable. | Std No
At Lab
MECHANICAL- DIMENSION (BASIC MEASURING INSTRUMENT, GAUGE ETC.) | Bore Gauge (For transmission) | Comparison by dial calibration tester | 0 to 1 mm
| Caliper (Vernier, Dial & Digital)(L.C. 0.01mm & coarser) | Comparison by using Caliper checker & slip gauge block | 0 to 600 mm
| Coating Thickness Meter (L.C. 0.001 mm) | Comparison by Thickness foils | 0 to 1 mm
| Comparator Stand For Flatness of Base | Comparison by Electronic comparator | 0 to 100 mm
| Comparator Stand For Flatness of Base | Comparison by Electronic comparator | 0 to 300 mm
| Depth Gauge (Vernier, Dial & Digital)(L.C. 0.01mm & coarser) | Comparison by Using Depth Micrometer Checker | 0 to 300mm
MECHANICAL- DIMENSION (BASIC MEASURING INSTRUMENT, GAUGE ETC.) | Depth Micrometer (L.C. 0.01mm & coarser) | Comparison by Depth Micrometer Checker | 0 to 300 mm
| Dial Calibration Tester (L.C. 0.0001mm) | Comparison by slip gauge block | 0 to 25 mm
| Dial Snap Gauge | Comparison by slip gauges | 1 mm to 200 mm
| Dial Thickness Gauge (L.C. 0.001 mm & coarser) | Comparison by Slip Gauge | 0 to 10 mm
| Electronic Probe With Indicator(L.C. 0.0001mm) | Comparison by Slip Gauge | 0 to 25 mm
| External Micrometer (L.C. 0.001mm & coarser) | Comparison by using Micrometer Checking set & Long Slip Gauge Block | 0 to 100 mm
| External Micrometer (L.C. 0.001mm & coarser) | Comparison by using Micrometer Checking set & Long Slip Gauge Block | 100 mm to 300 mm
| Feeler Gauge / Thickness foil | Comparison by electronic comparator | 0 to 1 mm
| Height Gauge (Vernier, Dial & Digital)(L.C. 0.01mm & coarser) | Comparison by using Caliper checker & Slip Gauge Block | 0 to 600 mm
| Inside Dial Caliper | Comparison by Slip gauges | 0 to 50 mm
MECHANICAL- DIMENSION (BASIC MEASURING INSTRUMENT, GAUGE ETC.) | Lever Dial (L.C. 0.001mm) | Comparison by dial calibration tester | 0 to 1 mm
| Measuring Pin | Comparison by Electronic Comparator & Slip Gauge | 0.5 mm to 25 mm
| Micrometer Setting Standard | Comparison by Electronic Comparator & Slip Gauge | 25 mm to 300mm
| Parallel Thread Plug Gauge | Comparison by FCDM,CSM & thread measuring wire | 2 mm to 300 mm
| Pistol Caliper (L.C. 0.10 mm) | Comparison by Slip Gauge | 0 to 50 mm
| Plain Plug Gauge / Width Gauge / Height Master / O.D. Gauge | Comparison by Electronic Comparator & Slip Gauge | 0 to 200 mm
| Plain Plug Gauge / Width Gauge / Height Master / O.D. Gauge | Comparison by Electronic Comparator & Slip Gauge | 200 mm to 400 mm
| Plain Snap Gauge | Comparison by slip gauges | 0.5 mm to 200 mm
| Plain Snap Gauge | Comparison by slip gauges | 200 mm to 400 mm
| Plunger Dial (L.C. 0.001mm & coarser) | Comparison by using dial calibration tester | 0 to 25 mm
| Taper Thread Plug Gauge | Comparison by FCDM,CSM & thread measuring wire | 2 mm to 100 mm
| Thread measuring wire | Comparison by Electronic Comparator & Slip Gauge | 0.17 mm to 6.35 mm
MECHANICAL- PRESSURE INDICATING DEVICES | Hydraulic Pressure - (Dial, Digital, Analogue gauges) | Comparison by Digital Pressure Gauge with hydraulic Test Pump as per DKD R6–1 | 0 to 70 bar
MECHANICAL- PRESSURE INDICATING DEVICES | Hydraulic Pressure - (Dial, Digital, Analogue gauges) | Comparison by Digital Pressure Gauge with hydraulic Test Pump as per DKD R6–1 | 0 to 700 bar
MECHANICAL- PRESSURE INDICATING DEVICES | Pneumatic Pressure - (Dial, Digital, Analogue gauges) | Comparison by Digital Pressure Gauge with pneumatic Test Pump as per DKD R6–1 | 0 to 7 bar
MECHANICAL- PRESSURE INDICATING DEVICES | Vacuum Pressure - (Dial, Digital, Analogue gauges) | Comparison by Digital Pressure Gauge with pneumatic Test Pump as per DKD R6–1 | -0.8 bar to 0 bar
MECHANICAL- TORQUE GENERATING DEVICES | Torque Wrench (Type I/Class B,C and Type II/Class A,B) Clockwise direction only | Comparison by torque transducer with indicator , IS16906:2018 | 2 Nm to 1000 Nm
At Site
MECHANICAL- DIMENSION (BASIC MEASURING INSTRUMENT, GAUGE ETC.) | Electronic Height Gauge(L.C. 0.0001 mm) | Comparison by long slip gauge blocks | 0 to 600 mm
| Surface Plate (Cast Iron & Granite) | Comparison by Spirit Level of sens. 0.01 mm/meter | 3000 mm to 3000 mm
MECHANICAL- PRESSURE INDICATING DEVICES | Hydraulic Pressure - (Dial, Digital, Analogue gauges) | Comparison by Digital Pressure Gauge with hydraulic Test Pump as per DKD R6–1 | 0 to 70 bar
| Hydraulic Pressure - (Dial, Digital, Analogue gauges) | Comparison by Digital Pressure Gauge with hydraulic Test Pump as per DKD R6–1 | 0 to 700 bar
| Pneumatic Pressure - (Dial, Digital, Analogue gauges) | Comparison by Digital Pressure Gauge with pneumatic Test Pump as per DKD R6–1 | 0 to 7 bar
| Vacuum Pressure - (Dial, Digital, Analogue gauges) | Comparison by Digital Pressure Gauge with pneumatic Test Pump as per DKD R6–1 | -0.8 bar to 0 bar
Thermal | Temp. indicator / controller | uni. Calibrator
| Glass Thermometer | wet bath, temp sensors
| Oven / furnace / autoclave / incubator / hot plate / water bath / Salt spray tester / Climatic chamber / Refrigerator / Deep freezer | temp sensors

## Thermal Master
Master instruments
1 | 6 1/2 DMM
2 | PT-100 Sensor
| Multifunction Calibrator |  | Ac
Programmable Timer
Universal Calibrator
Dry Bath
Standard RH Indicator With Sensor, Humidity Generated
Thermocouple K- Type
Thermohygrometer
Ultrasonic Flow meter
Techometer
Sound level meter
IS:6274 :method of calibrating liquid in glass thermometer
IS:7358 :specification of thermocouple
EURAMET cg-8:calibration of thermocouple
ASTM E220 13: standard test method for calibration of thermocouple by comparation method
DKD-R5-1 :calibration of resistance thermometer
DKD -R5-7 :calibration of climate chamber
API (6) standard 20H :calibration of industrial furnace
VDI /VDE 3511 Part 4.3-high temperature